A GUIDE TO THE INSTALLATION, OPERATION AND MAINTENANCE OF

NFX   DIRECT READING FLOWMETER

Operation.

Flow tube replacement.

Nixon Flowmeters Ltd, Badminton Close, Leckhampton, Cheltenham, Glos GL53 7BX

Tel- 01242 243006        Fax- 01242 222487           Email mail@nixonflowmeters.co.uk

# Text boxes

For best results the flowmeter mounting should be checked for vertical and any connections made leak checked.

The flowrate is read by viewing the top float edge against the graduated flowscale.

Optional flow control valves can be used to regulate or set the desired flow position.  For isolating the flow loop, a separate stop valve should also be fitted.

Providing the materials are compatible with the metered fluid and the operating limits observed, there should be no requirement for maintenance. In the event of breakage, a change of duty, or the need to clean a tube the following section may apply. In the event of accidental damage or loss the following spare parts are available:

Valve Assembly	-	Fine or Ultra-fine

Flow Tube		-	Specify fluid and units

Seal Set		-	Viton

4

5

4

6

2

3

Installation.

Low flow units are supplied complete assembled, ready for installation.

Higher flow units are supplied with the tube and float assembly separate for transit purposes. Assemble the tube and float by removing the bottom float stop and inserting the float. Replace the bottom stop and mount the flowtube assembly in the frame.

Angled connections are suitable for direct mounting to a panel. The panel cut-out details for fixing and connection are shown below.

Note: max operating pressure 20 bar non shock, temperature range -15(C to 120(C.

| mm | Compact | Standard | Long |
| --- | --- | --- | --- |
| a | 133 | 210 | 250 |
| b | 108 | 184 | 226 |
| c | 65 | 121 | 121 |
| Max. Panel Thickness 6mm (1/4") |

1

1

3